## Intro&Application
PART 3: TEST REPORT FORMAT FOR TYPE EVALUATION
INTRODUCTION
This Report Format applies for any kind of protein measuring instrument for grain (independent of its technology). It presents a standardized format for the results of the various tests and examinations, outlined in Part 2 clause 7 of OIML R xxx (201x), to which a type of protein measuring instrument for grain shall be submitted with a view to its approval based this OIML Recommendation.
It is recommended that all metrology services or laboratories evaluating and/or testing types of protein measuring instrument for grain according to OIML R xxx (201x), or to national or regional regulations based on that Recommendation, use this Report Format, directly or after translation into a language other than English or French. In case of a translation, it is highly recommended to leave the structure and the numbers of the clauses unchanged: in this case most of the contents are also understandable for those who can not read the language of the translation.
It is also recommended that this Report Format in English or in French (or in both languages) be transmitted by the country performing the tests to the relevant authorities of another country, when requested for issuing a national or regional type-approval.
In the practical application of the Report Format, a cover page shall be included by the Issuing Authority, and clauses 1 - 5 shall be included as a minimum.
APPLICABILITY OF THIS TEST REPORT FORMAT
In the framework of the OIML Certificate System for Measuring Instruments applicable to protein measuring instruments for grain in conformity with OIML R xxx (201x), use of this report format is mandatory, in French and/or in English with translation into the national languages of the countries issuing such certificates, if appropriate.
Implementation of this Report Format is informative with regard to the implementation of OIML Recommendation OIML R xxx (201x) in national regulations.
GUIDANCE FOR THE APPLICATION OF THIS TEST REPORT
Refer to Part 1 clause 2 of of OIML R xxx (201x) for definitions of terms, acronyms and symbols used.
The measurement unit is not always stated in the header row or column in the table.
The examiner is not expected to include the measurement unit with each recorded measurement result.
Percent by weight (abbreviated as' % w/w' or '%') is the measurement unit applicable for any values of the following:
- Protein content (PMB)
- MPE, error shift limits, maximum fault, etc.
- Basis moisture content (MB), and
- Actual "as is" moisture content is
The coloured fields should always be filled as appropriate.
Where the heading or label of a coloured field indicates "Pass/ fail" or "Pass/ fail/ NA", select the applicable option in the drop-down list that appears when the mouse is placed over the right side of the field.
NOTE: If it is impossible to enter the results in a computer, this Report Form may be printed and completed manually. In this case write "Pass", "fail" or "NA" in the coloured fields as appropriate.
In the Examination Checklist, the optional comments fields have a different colour to the mandatory fields (i.e. fields labelled "Yes/ No", "Pass/ fail" response or details of validation)
The mandatory fields relating to conditions to be specifiec by the national responsible body (within the limits suggested in OIML R xxx) have a different colour to the fields with the result from the assessment of the submitted type and documentation (i.e. fields labelled "Pass/ fail/ NA"). Select "NA" only when the requirement (or a variant of a requirement) is not adopted by the national responsible body.
In case a prescribed test or requirement is not relevant for the type of instrument to be tested or has not been adopted by the national responsible body, the reason why the test is omitted shall be clearly stated in the field “Comments”
In the test reports, the white fields with blue outline contain calculations and/or conditional formatting to highlight a suspect result.
THE EVALUATION REPORT
The format for the report is given on the following pages
To review or revise the formulae for calculated values in the file 20130409 - 4CD Grain Protein Measurement - Part 3 Test Report.xlsx, select Menu Option 'Review' and then select 'Unprotect Sheet'.

## Cover-IssuingAuthority
COVER PAGE BY THE ISSUING AUTHORITY (NATIONAL RESPONSIBLE BODY)

## Contents
CONTENTS OF PART 3: TEST REPORT FORMAT | Location
Introduction | Intro&Application
Applicability of this Test Report Format | Intro&Application
Guidance for the application of this Test Report | Intro&Application
The evaluation report | Intro&Application
1 Authority responsible for this Report | Authority&Synopsis
2 Synopsis of the results of the examination and tests | Authority&Synopsis
3 Summary of the results of the examination and tests
3.1 Examinations | Summary
3.2 Performance tests | Summary
4 General information about the application
4.1 Manufacturer | ApplicationGenInfo
4.2 Applicant | ApplicationGenInfo
4.3 Testing laboratories involved in the tests | ApplicationGenInfo
4.4 General information concerning the type | ApplicationGenInfo
4.5 Accessories supplied by the applicant | ApplicationGenInfo
4.6 Selection of sample instruments | ApplicationGenInfo
4.7 Adjustments and modifications | ApplicationGenInfo
4.8 Results of previous tests that were taken into account | ApplicationGenInfo
4.9 Additional information on the submitted type | ApplicationGenInfo
4.10 Documentation supplied by applicant | ApplicationGenInfo
4.11 Test equipment and grain used in type evaluation | ApplicationGenInfo
5 Examination details
5.1 Examination checklist – metrological and technical requirements | ExamDetails
5.2 Examination checklist – software-controlled devices and security | ExamDetails
5.3 Examination checklist – documentation requirements | ExamDetails
6 Type evaluation tests
6.1 Tests for time-related effects
6.1.1 Instrument warm-up time | WarmUp
6.1.2 Instrument drift and instability | DriftInstability
6.2 Tests for influence factor variations within the rated operating ranges
6.2.1 Instrument levelling | Levelling
6.2.2 Cold | Cold
6.2.3 Dry Heat | DryHeat
6.2.4 Damp heat steady state | DampHeat
6.2.5 AC mains voltage variations | VoltVariations
6.2.6 Variation in voltage supplied by external 12V and 24 V road vehicle batteries | BatteryV
6.3 Tests for disturbances
6.3.1 AC mains voltage dips and short interruptions | VoltDipsInterrupt
6.3.2 Electrical Bursts | BurstsMains
6.3.3 Radiated, radio-frequency, electromagnetic fields | RadiatedRF,EMFields
6.3.4 Conducted, radio-frequency fields | ConductedRF,EMFields
6.3.5 Electrostatic discharge | ElecDischarges
Mechanical shock | MechShock
6.3.6 Storage temperature | StorageTemp
6.3.7 Random vibration | RandVibration
6.4 Tests to assess the submitted calibrations
6.4.1 Accuracy and precision at reference conditions | Accuracy,r&R-Calibration
6.4.2 Sample temperature sensitivity | STS-Calibration

## Authority&Synopsis
1 AUTHORITY RESPONSIBLE FOR THIS REPORT
Organisation name:
Address:
Report number:
Application number:
Testing period: |  | to
Issue date of this Report:
Approver name:
Approver signature:
Stamp(s) if applicable:
2 SYNOPSIS OF THE RESULTS OF THE EXAMINATION AND TESTS
The tested samples of the type fulfils ALL the applicable requirements in OIML R xxx (201x):
Pass/fail
Comments:

## Summary
3 SUMMARY OF THE RESULTS OF THE EXAMINATION AND TESTS
3.1 Examinations
(*) Clause also contains recommendations in regard to specifications set by the national responsible body.
R xxx Part 1 - Metrological / technical / software requirements |  | Result | Location / ref
3* | Units of measurement
4 | Metrological requirements
4.1* | Applicable grains and PMB measuring ranges – specification
4.2* | Instrument environmental operating temperature – specification
4.3* | Grain sample operating temperature – specification
4.4* | Influence quantities – specification
4.5* | Maximum permissible error and other accuracy requirements
4.7 | Requirements for calibrations |  | 3.2 Performance Tests Summary and test reports
4.8 | Error due to variations in influence quantities
4.9 | Error due to changes in the instrument over time
5 | Technical requirements
5.1 | Checking facilities
5.2* | Manufacturer’s manual
5.3 | Markings
5.4* | Sample input and calibration selection
5.5 | Instrument construction
5.6 | Level indicating means
5.7 | Presentation of the measured value
6 | Requirements for software-controlled devices and security
6.1 | Specification of software requirements
6.2 | Electronic data storage and transmission
6.3 | Software documentation
6.4* | Provision for software and calibration security
9.4* | Maintenance and reconfiguration of the approved software
9.5 | In-field updates to grain calibrations
R xxx Part 2
7.1.2 | Documentation file
3.2 Performance Tests
R xxx Annex C: Type evaluation tests |  | Result | Location
C.4 | Tests for time related effects
C.4.1 | Instrument warm-up time
C.4.2 | Instrument drift and instability
C.5 | Tests for influence variations within the rated operating conditions
C.5.1 | Instrument levelling
C.5.2 | Cold
C.5.3 | Dry heat
C.5.4 | Damp heat
C.5.5 | AC mains voltage variation
C.5.6 | Variation in voltage supplied by external road vehicle batteries
C.6 | Tests for disturbances
C.6.1 | AC mains voltage dips, short interruptions and voltage variations
C.6.2 | Bursts (transients) on AC mains
C.6.3 | Radiated radiofrequency, electromagnetic susceptibility
C.6.4 | Conducted radio-frequency fields
C.6.5 | Electrostatic discharges
C.6.6 | Storage temperature (extreme shipping conditions)
C.6.7 | Random vibration
(continued next page)
R xxx Annex C: Type evaluation tests (continued) |  | Result | Location
C.7 | Assessment of calibrations in the submitted type
GT1 | *
GT2 | *
GT3 | *
GT4 | *
GT*
C.7.1 | Accuracy and precision at reference conditions
GT1 | *
GT2 | *
GT3 | *
GT4 | *
GT*
* Insert Calibration name and version number
Note on J9: 201210 Leida Queddeng (NMIA): Example - Leave "Location / ref" blank. Print report to PDF and then insert respective page number in field and re-printing.

## ApplicationGenInfo
4 GENERAL INFORMATION ABOUT THE APPLICATION
4.1 Manufacturer
Name:
Address:
4.2 Applicant
Organisation:
Contact name:
Address:
Email address:
Phone: |  |  |  | Fax:
Application date: |  |  |  | Ref number:
Applicant is authorised by the manufacturer (documented evidence) | [6 empty] | Yes/ no
Comments:
4.3 Testing laboratories involved in the tests
Complete this form for each test laboratory -
Organisation name:
Address:
Application number:
Tests by this laboratory:
Testing period: |  |  |  |  | to
Name(s) of test engineer(s):
Laboratory accredited by:
Accreditation number: |  |  |  |  | Expiry:
Accreditation includes R xxx: |  |  | Yes/ no |  | Edition:
Details of relevant peer assessement or assessment by other means:
Details, if any tests have been performed at another location than the laboratory premises:
Responsible person - name:
Date signed:
Signature:
Stamp(s) if applicable:
Comments:
4.4 General information concerning type
4.4.1 Description of the instrument (key technical characteristics and intended applications)
e.g. benchtop near infrared (NIR) protein measuring instrument with calibrations for wholegrain wheat and barley
4.4.2 Information indicated on the instrument
Manufacturer trademark:
Year of manufacture:
Type designation:
Model number(s) (if applic):
Electrical power markings:
Software ID (if applic):
Other descriptor/marking:
Comments:
4.4.3 Information on sample units
Serial number |  | Model number |  |  | Manufacture mode (prototype / production) |  | Year
Add additional rows if over three sample units are submitted.
Comments:
4.4.4 Relevant external/internal photographs taken during the examination and tests
4.5 Accessories supplied by the applicant
Batteries (if applicable): |  | Type |  | Vnom |  | Number required
Data printer (if applic):
External data storage (if applic):
Cables:
Other accesories:
4.6 Information on sample instruments
In case the tests and evaluation are valid for more versions, give full details of the ypes, versions,
measuring ranges, etc.:
Justification for the selection of the sample unit(s):
4.7 Adjustments and modifications
Adjustments, modifications and repairs made to the sample unit(s) during the testing:
4.8 Results of previous tests that were taken into account
Details:
4.9 Additional information concerning type
4.9.1 Instrument limitations of use
Calibration dependent
Sample PMB content (%) |  | Min: |  | Max:
Sample moisture content (%) |  | Min: |  | Max:
Operating temperature (°C) |  | Min: |  | Max:
Grain sample temperature (°C) |  | Min: |  | Max:
Maximum ΔT (°C) |  | or ΔTC,max |  | & ΔTH,max
If limits are dependent on the grain calibration (i.e. answer is 'Yes') specify values/range in 4.9.2 instead.
Comments:
4.9.2 Information on the submitted calibrations
Calibration principle:
Calibration number |  | GT1 |  |  | GT2 |  | etc.*
Calibration name:
Version number:
Displayed name:
Date submitted:
Calibration limitations of use -
Applicable type(s) of grain:
PMB measuring range:
Min/max moisture content:
TC,sample and TH,sample:
ΔTmax or ΔTCmax and ΔTHmax:
Regression information -
Approx number of data points:
Data sources, date range:
Reference method(s):
Other validation result:
(e.g. SD, SEP)
Default bias (if applic):
Default slope (if applic):
Other characteristic:
*Copy table into additional pages if more than two calibrations are submitted for examination.
Comments:
4.9.3 Additional information (e.g. connection equipment, interfaces, etc.)
4.10 Documentation supplied by applicant
Date received | Document title and/or reference number |  |  |  | Description (include version number if applicable)
Insert additional rows as required.
4.11 Test equipment and grain used in type evaluation
4.11.1 Test equipment
Instrument/ equipment | Make/ model | Serial # | Parameter applied/ measured |  | Calibrated range(s) | Test(s) used
Insert additional rows as required.
Details e.g. ~ equipment set-up for ESD and EMS tests
~ details of simulations
~ confidence intervals for uncertainty estimations
4.11.2 Grain test samples
Information about the grain reference materials (RMs) used in tests:
Reference method used to generate whole-grain certified reference materials (CRMs):

## ExamDetails
5 EXAMINATION DETAILS
5.1 Examination checklist - metrological and technical requirements
(*) Clause contains requirements to be addressed by the national responsible body.
Checks on the requirements within R xxx Part 1 |  |  |  | Findings
3 | Units of measurement
3.1 | The measurement of the protein content in a grain sample is expressed in percentage by mass (% w/w). The percentage symbol alone (%) is also permitted.
Pass/ fail
3.2 | For each type of grain, the measured protein content is expressed at one basis moisture content (MB). The scaling of the protein content at the actual “as is” moisture content (PM) to the protein content at the basis moisture content (PMB) is in accordance with Equation 1.
Pass/ fail
3.3* | The national responsible body has specified the basis moisture content (MB) for expressing the protein content of each grain type. Examiner to append and/or reference the national specification.
Yes/ no
4 | Metrological requirements
4.1 | Applicable grains and PMB measuring ranges – specification
4.1.1* | The national responsible body has specify commercially important PMB ranges for the grains on which PMB measurements are subject to national approval. R xxx Table 3 contains examples of the types of grains that are commercially important in certain countries/ regions. Examiner to append and/or reference the national specification for the PMB measurement ranges on commercially important grains.
Yes/ no
4.1.2 | The grain types and the corresponding PMB measuring range that can be analysed by the submitted type of instrument has been declared by the manufacturer. The PMB measuring range includes the range specified by the national responsible body for the grain type. Examiner to append and/or reference the manufacturer’s specification.
Pass/ fail
4.2 | Instrument environmental operating temperature – specification
4.2.1(a)* | Examiner to indicate whether a limited range for the environmental operating temperature is permitted. If the answer is ‘No’ the temperatures applied during type testing should include all the possible environmental temperatures in that particular country.
Yes/ no
4.2.1(b) | The following only applies if the previous answer is 'No'- The submitted type provides accurate PMB measurements in all the environmental temperatures possible in the country/region. Examiner response should be based on the relevant performance test(s).
Pass/ fail/ NA
4.2.2* | The following only applies if the answer to 4.2.1 is ‘Yes’- The national responsible body has specified the range of ambient temperatures (TC to TH) in which the instrument can be used to take PMB measurements for commercial purposes. The temperature range TC to TH includes 10 ºC to 30 ºC. Examiner to append the national specification or indicate the values for TC and TH. |  |  | TC = TH =
Yes/ no/ NA
4.2.3(a) | The following only applies if the answers to 4.2.1(a) and 4.2.2 are both ‘Yes’- Examiner to indicate whether the manufacturer has specified a wider ambient temperature range than specified by the national responsible body.
Yes/ no/ NA
4.2.3(b)* | If the previous answer is ‘Yes’, the manufacturer may request for the wider range to be adopted for their type evaluation. Examiner to indicate whether the manufacturer specification has been adopted as TC to TH for this particular type evaluation. If the answer is ‘Yes’, append the manufacturer specification and indicate the revised values for TC to TH. |  |  | Values adopted for approval of this type: TC = TH =
Yes/ no/ NA
Continued - Checks on the requirements within R xxx Part 1 |  |  |  | Findings
4.3 | Grain sample operating temperature – specification
4.3.1 | Specification of the sample temperature range
4.3.1.1 | Examiner to indicate whether a limited range for the grain sample temperature is permitted. If the answer is ‘No’ the temperature of grain samples during type testing should encompass the range 2 ºC to 45 ºC.
(a)*
Yes/ no
4.3.1.1 | The following only applies if the previous answer is 'No'- The submitted type provides accurate PMB measurements on samples that are between 2 ºC to 45 ºC in temperature. Examiner response should be based on the relevant performance test.
(b)
Pass/ fail/ NA
4.3.1.2* | The following only applies if the answer to 4.3.1.1 is ‘Yes’- The national responsible body has specified the range of sample temperatures (TC,sample to TH,sample) in which the instrument can be used to take PMB measurements for commercial purposes. The temperature range TC,sample to TH,sample includes 10 ºC to 30 ºC. Examiner to append national specification to report or indicate the values for TC,sample and TH,sample for each grain type. |  |  | TC,sample = TH,sample =
Yes/ no/ NA
4.3.1.3 | The following only applies if the answers to 4.3.1.1(a) and 4.3.1.2 are ‘Yes’- Examiner to indicate whether the manufacturer has specified a wider sample temperature range than specified by the national responsible body.
(a)
Yes/ no/ NA
4.3.1.3 | If the previous answer is ‘Yes’, the manufacturer may request for the wider range to be adopted for their type evaluation. Examiner to indicate whether the manufacturer specification has been adopted as TC,sample to TH,sample for this particular type evaluation. If the answer is ‘Yes’, append the manufacturer specification and indicate the revised values for TC,sample to TH,sample for each grain type. |  |  | Values adopted for approval of this type: TC,sample = TH,sample =
(b)*
Yes/ no/ NA
4.3.2 | Specification of the sample and instrument maximum temperature differential (ΔTmax)
4.3.2.1 | Examiner to indicate whether a maximum limit for ΔT (i.e. ΔTmax) is permitted. If the answer is ‘No’ the following values shall be applied during type testing: ΔTC,max = Tref – TC,sample and ΔTH,max = TH,sample – Tref.
(a)*
Yes/ no
4.3.2.1 | The following only applies if the previous answer is 'No'- The instrument at reference temperature (Tref), provides accurate PMB measurements on samples within the range TC,sample to TH,sample regardless of the magnitude of the sample and instrument temperature differential (ΔT). Examiner response should be based on the relevant performance test.
(b)
Pass/ fail/ NA
4.3.2.2* | The following only applies if the answer to 4.3.2.1 is ‘Yes’- The national responsible body has specified a ΔTmax in which the instrument can be used to take PMB measurements for commercial purposes. The value of ΔTmax (or ΔTC,max and ΔTH,max, if unequal about Tref) is at least 10 ºC. Examiner to append national specification to report or indicate the values for ΔTmax (or ΔTC,max and ΔTH,max) for each grain type. |  |  | ΔTmax = ΔTC,max (if ≠ ΔTH,max): ΔTH,max (if ≠ ΔTC,max):
Yes/ no/ NA
4.3.2.3 | The following only applies if the answers to 4.3.2.1(a) and 4.3.2.2 are ‘Yes’- Examiner to indicate whether the manufacturer has specified a larger maximum differential than ΔTmax specified by the national responsible body.
(a)
Yes/ no/ NA
4.3.2.3 | If the previous answer is ‘Yes’, the manufacturer may request for the wider range to be adopted for their type evaluation. Examiner to indicate whether indicate whether the manufacturer specification has been adopted as ΔTmax for this particular type evaluation. If the answer is ‘Yes’, append the manufacturer specification and indicate the revised values for ΔTmax (or ΔTC,max and ΔTH,max) for each grain type. |  |  | Values adopted for approval of this type: ΔTmax = ΔTC,max (if ≠ ΔTH,max): ΔTH,max (if ≠ ΔTC,max):
(b)*
Yes/ no/ NA
4.3.3 | Provisions in absence of a manufacturer-specified sample temperature range
(a) | Examiner to indicate whether the manufacturer has declared a sample temperature operating range or a maximum allowable ΔT. If the answer is ‘No’ append and/or provide reference the relevant national testing and/or operating procedures.
Yes/ no/ NA
Continued - Checks on the requirements within R xxx Part 1 |  |  |  | Findings
4.3.3 | The following only applies if the previous answer is 'No'- Testing and/or operating procedures are in place to ensure that accurate PMB measurements are possible within the testing ranges specified for type evaluation (i.e. TC,sample to TH,sample and ΔTmax).
(b)*
Yes/ no/ NA
4.4 | Influence quantities – specification
4.4.1* | Rated operating ranges for influence factors specified by the national responsible body conform to the international standard. Examiner to indicate reasons for deviations.
Yes/ no
4.4.2* | Disturbance tests for electronic instruments specified by the national responsible body conform to the international standard. Examiner to indicate reasons for deviations.
Yes/ no
4.5 | Maximum permissible error (MPE) and other accuracy requirements
4.5.2* | Accuracy requirements specified by the national responsible body for each grain type conform to the international standard. Examiner to indicate reasons for deviations.
Yes/ no
4.5.3* | The MPE for each grain type has been scaled in accordance with the nationally-specified MB and the recommended rounding method performed.
Yes/ no
4.5.4* | Reference method specified for the national responsible body for PMB measurements is based on an international standard. Examiner to indicate reasons for deviations and to append and/or reference the test procedure.
Yes/ no
5 | Technical requirements
5.1 | Checking facilities
5.1.1 | Suppression of PMB measured values in the event of a significant fault. There is provision for the instrument to automatically and clearly indicate when a significant fault has occurred by an appropriate error message, unambiguous warning or blanking the display.
Pass/ fail
5.1.2 | Suppression of PMB measured values outside of operating ranges
5.1.2.1 | The operator is not required to judge the precise ambient temperature and the temperature of the sample required in order to make an accurate measurement. The instrument automatically and clearly indicates when a type-approved operating range is exceeded by an appropriate error message, unambiguous warning or blanking the display.
Pass/ fail
5.1.2.2 | The instrument shall automatically prevent further measurements as long as the respective influence factor or sample characteristic remains outside the type-approved ranges.
Pass/ fail
5.2 | Manufacturer's manual
5.2.1 | There is a manual to be provided with each protein measuring instrument, that describes the installation, operation, and routine maintenance of the instrument and its accessories.
Pass/ fail
5.2.2 | The manual must also include the following information: (a) name and address of the manufacturer; (b) the type of the instrument with which it is intended to be used; (c) date of issue; (d) the types of grain for which the instrument is designed to be used within the scope of national requirements; (e) the limitations of use (refer to OIML R xxx Part 1 for details).
Pass/ fail
5.2.3(a)* | The national responsible body has indicated acceptable language(s). Examiner to note the accepted language(s) if 'Yes' or the official language(s).
Yes/ no
5.2.3(b) | The user/owner manual is supplied in all the accepted or official language(s) and the translations appear to be accurate.
Pass/ fail
5.3 | Markings
5.3.1 | Instrument is clearly and permanently marked with the following: (a) manufacturer’s name or mark; (b) model designation; and (c) serial number given by the manufacturer. Provision is made for application of a type approval mark.
Pass/ fail
Continued - Checks on the requirements within R xxx Part 1 |  |  |  | Findings
5.3.2 | Markings are grouped together in a clearly visible location, either on a permanently attached nameplate or on part of the instrument. The required information is readily observable without disassembly.
Pass/ fail
5.4 | Sample input and calibration selection
5.4.1 | Provisions in absence of a manufacturer-specified sample temperature range
5.4.1.1 | On instruments with different calibrations, the user is able to select the calibration applicable for the sample to be analysed.
Pass/ fail
5.4.1.2 | The selection of the calibration on the user interface is unambiguous and visible to all parties present, i.e. during the measurement it can be verified that the selected (displayed) calibration corresponds with the sample analysed.
Pass/ fail
5.4.2 | Sampling and minimum sample size
5.4.2.1 | The operator shall not be required to judge the precise volume or weight required by the instrument to make an accurate PMB measurement.
Pass/ fail
5.4.2.2 | Examiner to indicate whether sample smaller than the international recommendation is permitted by the national responsible body. If the answer is 'Yes', examiner to indicate the smallest sample size permitted.
(a)*
Yes/ no
5.4.2.2 | The following only applies if the previous answer is 'No'- The size of the sample analysed is no less than 100 g or 400 kernels or seeds which ever is smaller.
(b)
Pass/ fail/ NA
5.4.2.3 | The national responsible body has specified minimum guidelines for the sampling of cereals for testing based on international standards. Examiner to indicate reasons for deviations.
Pass/ fail
5.5 | Instrument construction
5.5.1 | Nothing observed in the design and construction of the instrument and accessory equipment should make it prone to inaccuracy, malfunction and fraud under normal service conditions.
Pass/ fail
5.5.2 | Day to day forces on the parts of the instruments shall not affect the accuracy of measurements.
Pass/ fail
5.5.3 | The instrument housing protects the main components from dust and moisture.
Pass/ fail
5.5.4 | The following only applies if the instrument analyses ground/milled samples- The manufacturer of the grain protein measuring instrument has designated the type(s) of mill to be used and included a unit with the submission so its suitability for the measurement process be assessed during type evaluation.
Pass/ fail/ NA
5.6 | Level indicating means
5.6.1(a) | The performance of the instrument is changed by an amount greater than the maximum error shift when the instrument is moved from a level position into a position that is out of level in any upright direction by up to 5% (approx. 3º). Examiner response should be based on the relevant performance test.
Yes/ no
5.6.1(b) | The following only applies if the previous answer is 'Yes'- The instrument is equipped with a level indicator and level adjustment means.
Pass/ fail/ NA
5.6.2 | The following only applies if the answer to 5.6.1(a) is 'Yes'- The level indicating means is readable without any instrument disassembly.
Pass/ fail/ NA
5.7 | Presentation of the measured value
5.7.1 | The instrument is equipped with a digital indicating element which does not display any protein content values before the end of the measurement cycle.
Pass/ fail
5.7.2 | Measurement results are displayed as percent protein by mass (%) the MB. Subdivisions of this unit are in terms of decimal subdivisions (not fractions).
Pass/ fail
5.7.3(a) | The display on the type allows the protein content value to be determined with a resolution of at least 0.1% PMB.
Pass/ fail
5.7.3(b) | Units submitted for type evaluation permit 0.01% PMB resolution. |  |  |  | Pass/ fail
5.7.4(a) | The type is multi-constituent measuring instrument (e.g. it measures grain moisture content in addition to grain protein content)
Yes/ no
5.7.4(b) | The following only applies if the previous answer is 'Yes'- Appropriate labels are displayed or recorded to make it clear which constituent is associated the displayed or recorded measured values.
Pass/ fail/ NA
5.7.5 | The height for the digits used to display protein content is at least 10 mm. Numbers and symbols of units are presented in accordance with OIML D 2.
Pass/ fail
5.2 Examination checklist - requirements for software-controlled devices and security
Checks on the requirements within R xxx Part 1
6.1 | Specification of the software requirements
6.1.1 | For instruments and modules operated by software, the manufacturer has described or declared how the software is implemented within the instrument or module. Examiner to indicate whether the software is embedded or on an universal computer system.
| Validation method: | AD | Provide validation details below- |  | Pass/ fail
6.1.2 | The legally relevant software is clearly identifiable via a unique software version and a checksum. Examiner to indicate whether the software version and the checksum is displayed or printed out on command during operational mode, or displayed during the start-up procedure.
| Validation method(s): | AD + VFTSw | Provide validation details below- |  | Pass/ fail
6.1.3 | The legally relevant measuring algorithms and functions are appropriate and functionally correct. Examiner to conduct further examinations and tests if necessary.
| Validation method(s): | AD + VFTM | Provide validation details below- |  | Pass/ fail
6.1.4 | Optional check during type evaluation- Conformity of the legally relevant software (in the submitted units) to that in the approved type has been verified at level (b) described in D 31 clause 5.2.5. This may be confirmed by noting the identity of parts of the legally relevant source code, and for the rest of the legally relevant software, the identity of the legally relevant functions described in the documentation.
Pass/ fail/ NA
6.1.5 | There is provision to make further measurements impossible when a significant fault is detected.
| Validation method(s): | AD + VFTSw | Provide validation details below- |  | Pass/ fail
6.1.6(a) | The software of the instrument is separated in legally relevant part and non-relevant parts.
Yes/ no
6.1.6(b) | The following only applies if the previous answer is 'Yes'- The requirements of D 31 clause 5.2.1.2 regarding Separation of software parts have been fulfilled.
| Validation method(s): | AD + VFTSw | Provide validation details below- |  | Pass/ fail/ NA
6.1.7(a) | The instrument uses an internal or external universal computer. |  |  |  | Yes/ no
6.1.7(b)* | Use of the protein measuring instrument in an open network is permitted. |  |  |  | Yes/ no
6.1.7{c)* | Software changes on a verified instrument by a 'Traced Update' is permitted. |  |  |  | Yes/ no
6.1.7(d) | The following applies if the answer to any of the above 6.1.7(a) - 6.1.7{c) is 'Yes'- The legally relevant software can be operated only in the environment specified for its correct functioning. If necessary to secure the correct functioning of the legally relevant software, the operating system has been fixed to a defined invariant configuration.
| Validation method(s): | AD + VFTSw | Provide validation details below- |  | Pass/ fail/ NA
6.1.8(a)* | Regarding industry trading practices in the country where approval is sought: Legally relevant measurement data may leave the measuring instrument and be stored or transmitted before they are used for commercial purposes. The requirements in clauses 6.2.1 - 6.2.5 only apply if the answer is 'Yes'.
Yes/ no
6.1.8(b)* | The national responsible body requires instruments to be equipped with an internal recording element and/or a communication interface that permits interfacing with an external recording element. Examiner to append and/or reference details of this requirement.
Yes/ no
6.1.8(c) | The following only applies if the answer to 6.1.8(b) is 'Yes'- The instrument is equipped with the required recording element(s).
Pass/ fail/ NA
6.1.8(d)* | The following only applies if the answer 6.1.8(a) or 6.1.8(b) is 'Yes'- Data storage and/or transmission is allowed to take place in an insecure environment. NOTE: Ensure consistency with the answer in 6.1.7(b).
Yes/ no/ NA
Continued - Checks on the requirements within R xxx Part 1 |  |  |  | Findings
6.2.1(e)* | The following only applies if the answer to 6.1.8(d) is 'Yes'- The national responsible body requires data to be protected by cryptographic means where instruments can be used in an open network.
Yes/ no/ NA
6.2 | Electronic data storage and transmission
6.2.1 | The measurement value stored or transmitted is accompanied by all relevant information necessary for future legally relevant use. Examiner to ensure that the measurement records include: test sample identifier, measurement date and time, unique identification of the instrument, grain type, PMB results and units as displayed, calibration version ID, error messages and constituent labels (on multi-constituent meters).
| Validation method(s): | AD + VFTSw | Provide validation details below- |  | Pass/ fail/ NA
6.2.2(a) | The following only applies if the answer to 6.2.1(d) is 'Yes'- The data is protected by software means to guarantee the authenticity and integrity. There is provision for the data to be discarded or marked unusable if an irregularity is detected.
| Validation method(s): | AD + VFTSw | Provide validation details below- |  | Pass/ fail/ NA
6.2.2(b) | The following only applies if the answer to 6.2.1(e) is 'Yes'- Data is protected by cryptographic means. A seal is broken if a confidential key is input or read.
| Validation method(s): | AD + VFTSw + SMT | Provide validation details below- |  | Pass/ fail/ NA
6.2.3 | The measurement data is stored automatically when the measurement is concluded. No protein content values are recorded before the end of the measurement cycle. The storage device has sufficient permanency to ensure that data are not corrupted. There is sufficient memory for storage of the required measurement data to be used at a later time.
| Validation method(s): | AD + VFTSw | Provide validation details below- |  | Pass/ fail/ NA
6.2.4 | There is provision so that the measurement is not inadmissibly influenced by a transmission delay.
| Validation method(s): | AD + VFTSw | Provide validation details below- |  | Pass/ fail/ NA
6.2.5 | There is provision to ensure measurement data is not lost if a transmission interruption occurs because the network services become unavailable.
| Validation method(s): | AD + VFTSw | Provide validation details below- |  | Pass/ fail/ NA
6.4 | Provision for software and calibration security
6.4.1 | Sealing: Provision has been made for appropriate sealing by mechanical, electronic and/or cryptographic means, making any change that affects the metrological integrity of the instrument impossible or evident. Calibrations, zero-setting and test point adjustments are sealed.
| Validation method(s): | AD + VFTSw | Provide validation details below- |  | Pass/ fail
6.4.2 | Safeguards against fraudulent use
6.4.2.1 | The legally relevant software has been secured against unauthorised modification, loading or changes by swapping of the memory device. If the instrument has an operating system or an option to load software, additional means to mechanical sealing have been considered.
6.4.2.2 | Only clearly documented functions are allowed to be activated by the user interface, which have been realised in such a way that it does not facilitate fraudulent use.
6.4.2.3 | Parameters that fix the legally relevant characteristics of the measuring instrument shall be secured against unauthorised modification. The current parameter settings can be printed or displayed on demand.
| Validation method(s): | AD + VFTSw | Provide validation details below- |  | Pass/ fail
6.4.2.4 | The national responsible body restricts access to device-specific parameters. Examiner to append and/or reference details of any access restrictions.
(a)* |  |  |  |  | Yes/ no
6.4.2.4(b) | The instrument complies with access restrictions to adjustable parameters. |  |  |  | Pass/ fail/ NA
Checks on the requirements within R xxx Parts 1 & 2 |  |  |  | Findings
6.3 | Software documentation
The submitted software documentation is complete. Examiner to indicate inclusion of the following information with the submission-
Pass/ fail
Description of the legally relevant software, incorporating how the requirements are met
Yes/ no
Description of the operating system security
Yes/ no
Description of the software sealing method(s)
Yes/ no
Overview of the system hardware, highlighting any hardware components that are deemed legally relevant or performing legally relevant functions
Yes/ no
Description of the accuracy of the algorithms
Yes/ no
Declaration of the hardware and software environment, including minimum resources and configuration necessary for correct functioning of the instrument
Yes/ no
Description of the user interface, menus and dialogues
Yes/ no
Description of the software identification which has to be clearly assigned to the legally relevant functions
Yes/ no
Clear instructions on how to check the actual software identification against the reference number as listed in the type approval certificate
Yes/ no
List of commands of each hardware interface of the measuring instrument/ electronic device/ sub-assembly
Yes/ no
List of durability errors that are detected by the software
Yes/ no
Description of data sets stored or transmitted
Yes/ no
List of significant faults that are detected and a description of the detecting algorithm
Yes/ no
Operating manual which clearly identifies all operational controls, indications, and features
Yes/ no
9.4 | Maintenance and reconfiguration of the approved software
9.4(a)* | The national responsible body allows for software changes on the instrument by a Traced Update instead or in addition to Verified Updates. NOTE: Ensure consistency with the response to 6.1.7{c).
Yes/ no
9.4(b) | The following only applies if the previous answer is 'Yes'- Implementation of the software fulfils the requirements for Traced Updates in D 31 clause 5.2.6.2.
| Validation method(s): | AD + VFTSw | Provide validation details below- |  | Pass/ fail/ NA
9.5 | In-field updates to grain calibrations
9.5.1 | Seasonal variability of crops and inevitability calibration updates
9.5.1(a)* | The national responsible body requires that grain PMB calibrations included with the instrument can be updated to accommodate for seasonal variations. The requirements in 9.5.2 and 9.5.3 only apply if the answer is 'Yes'.
Yes/ no
9.5.1(b) | The following only applies if the previous answer is 'Yes'- The submitted type complies with any requirements for adjustable calibrations.
Pass/ fail/ NA
9.5.1(c)* | The national responsible body requires retention of data collected during the current and/or recent years that are used for calibration adjustments.
Yes/ no
9.5.2 | Calibration version
The calibration constants that are adjustable and unique calibration names, or calibration version numbers can be displayed and printed on demand.
Pass/ fail/ NA
9.5.3 | Security of calibrations and reverification
There is provision for only authorised persons to change grain calibrations. The security level for updating calibrations is the same as for software installation. There is provision for an error message to be automatically displayed if calibration constants are electronically corrupted. There is provision to make further measurements impossible.
| Validation method(s): | AD + VFTSw | Provide validation details below- |  | Pass/ fail
5.3 Examination checklist - documentation requirements
Checks on the requirements within R xxx Part 2 |  |  |  | Findings
7.1.2 | Documentation file
The submitted documentation file is complete. Examiner to indicate inclusion of the following information with the submission-
Pass/ fail
General description of the instrument
Yes/ no
General characteristics of the instrument
Yes/ no
Drawings of general arrangement and details of metrological interest
Yes/ no
Description of calibrations submitted for approval
Yes/ no
Descriptions and characteristic data for all devices and sub-assemblies of the instrument
Yes/ no
Technical description, drawings and plans of devices, sub-assemblies
Yes/ no
Declarations of the manufacturer
Yes/ no
Samples of all intended print-outs
Yes/ no
Information concerning special cases
Yes/ no
Results of tests performed by the manufacturer using protocols from Parts 2 and 3
Yes/ no
Certificates of other type approvals or separate tests
Yes/ no
Drawing or photo of the instrument
Yes/ no
Manufacturer’s manual according to OIML R xxx clause 5.2. NOTE: Ensure consistency with the responses in 5.2.1, 5.2.2 and 5.2.3(b).
Yes/ no
Software documentation according to OIML R xxx clause 6.3
Yes/ no
Other evidence to support the assumption that the design and characteristics of the type comply with the requirements of R xxx
Yes/ no

## WarmUp
6 TYPE EVALUATION TESTS
6.1 Tests for time-related effects
6.1.1 Instrument warm-up time [ref. OIML R xxx Annex C clause C.4.1]
Observer: | [5 empty] | General comments on test:
Type/ application #:
Instrument 1 ID:
Instrument 2 ID:
|  | Not warm | Warm
Ambient temp: |  |  |  | °C
Ambient RH: |  |  |  | %
Date commenced: |  |  |  | ddmmyyyy
Time commenced: |  |  |  | hh:mm
Displayed | Error shift limits |  | Sample info |  | Nominal values (%)
name
GT | Min | Max | ID |  | PMB | moisture
Instrument | Mean PMB |  | Error shift |  | RESULTS SUMMARY (Pass / fail)
| Not warm | Warm | Warm |  | Comments |  | Instrument | Type
1
2
Raw data entry
| PMB (Not warmed-up) |  |  |  | PMB (Warm-up time elapsed)
| Instr. 1 | Instr. 2 | Comments |  | Instr. 1 | Instr. 2 | Comments
1
2
3
4
5
6

## DriftInstability
6.1.2 Instrument drift and instability [ref. OIML R xxx Annex C clause C.4.2]
Observer: | [5 empty] | General comments on test:
Type/ application #:
Instrument 1 ID:
7.1.2 Instrument 2 ID:
|  | Before testing received | After 4-6 wks
Ambient temp: |  |  |  | °C
Ambient RH: |  |  |  | %
Date commenced: |  |  |  | ddmmyyyy
Time commenced: |  |  |  | hh:mm
Displayed | Error shift limits |  | Sample info |  |  | Nominal values (%)
name
GT | Min | Max | number | ID |  | PMB | moisture
1
2
3
Sample number | Instrument | Mean PMB |  | Error shift |  | RESULTS SUMMARY (Pass / fail)
|  | Start | 4 - 6 wks | After 4 - 6 weeks |  | Comments |  | Instrument | Type
1 | 1
2
3
1 | 2
2
3
Raw data entry
Sample number | PMB (Before testing, except warm-up test) |  |  |  | PMB (After 4 - 6 weeks)
| Instr. 1 | Instr. 2 | Comments |  | Instr. 1 | Instr. 2 | Comments
1
2
3

## Levelling
6.2 Tests for influence factor variations within the rated operating ranges
6.2.1 Instrument levelling [ref. OIML R xxx Annex C clause C.5.1]
Observer: | [6 empty] | Displayed | Error shift limits |  | Sample info |  | Nominal values (%)
Type/ application #: | [6 empty] | name
Instrument 1 ID: | [6 empty] | GT | Min | Max | ID |  | PMB | moisture
Instrument 2 ID:
|  | Start - ref | Influence |  |  | Recovery |  |  | General comments on test (e.g. tilt orientation):
|  | Level | 1st tilt | 2nd tilt | * | Level
Tilt angle: | [6 empty] | ° (angluar degrees)
Ambient temp: | [6 empty] | °C
Ambient RH: | [6 empty] | %
Date commenced: | [6 empty] | ddmmyyyy
Time commenced: | [6 empty] | hh:mm
* Add other tilt orientations if required
Instrument | Mean PMB |  |  |  |  | Error shift (at tilt orientation) |  |  |  | RESULTS SUMMARY (Pass / fail)
| Start - ref | 1st | 2nd | * | Recovery | 1st | 2nd | * | Recovery | Comments |  | Instrument | Type
1
2
Raw data entry
PMB (Start -ref) |  |  | PMB (1st orientation) |  |  | PMB (2nd orientation) |  |  | PMB (*) |  |  | PMB (Recovery)
Instr. 1 | Instr. 2 | Comments | Instr. 1 | Instr. 2 | Comments | Instr. 1 | Instr. 2 | Comments | Instr. 1 | Instr. 2 | Comments | Instr. 1 | Instr. 2 | Comments

## Cold
6.2.2 Cold [ref. OIML R xxx Annex C clause C.5.1]
Observer: | [5 empty] | General comments on test:
Type/ application #:
Instrument 1 ID:
Instrument 2 ID:
Spare instrument ID:
|  | Starf ref | Cold | Recovery
EUT ambient temp: |  |  |  |  | °C
EUT ambient RH: |  |  |  |  | %
Sample temp: |  |  |  |  | °C
Spare ambient temp: |  |  |  |  | °C
Spare ambient RH: |  |  |  |  | %
Date commenced: |  |  |  |  | ddmmyyyy
Time commenced: |  |  |  |  | hh:mm
Displayed | Error shift limits |  | Sample info |  |  | Nominal values (%) |  | Sample stability | Correct for recov
name
GT | Min | Max | number | ID |  | PMB | moisture | Δ PMB | Yes/ No
1
2
3
Sample | Instrument | Error shift |  |  | RESULTS SUMMARY (Pass/fail)
number |  | Cold | Recovery | Corrected | Comments |  | Instrument | Type
1 | 1
2
3
1 | 2
2
3
Sample | Mean PMB (Start - ref) |  |  | Mean PMB (Cold) |  | Mean PMB (Recovery)
number | Instr. 1 | Instr. 2 | Spare-ref | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | Spare-ref
1
2
3
Sample | PMB (Start - ref) |  |  | PMB (Cold) |  | PMB (Recovery) |  |  | Comments
number | Instr. 1 | Instr. 2 | Spare-ref | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | Spare-ref
1
2
3
Comments

## DryHeat
6.2.3 Dry heat [ref. OIML R xxx Annex C clause C.5.3]
Observer: | [5 empty] | General comments on test:
Type/ application #:
Instrument 1 ID:
Instrument 2 ID:
Spare instrument ID:
|  | Starf ref | Dry heat | Recovery
EUT ambient temp: |  |  |  |  | °C
EUT ambient RH: |  |  |  |  | %
Sample temp: |  |  |  |  | °C
Spare ambient temp: |  |  |  |  | °C
Spare ambient RH: |  |  |  |  | %
Date commenced: |  |  |  |  | ddmmyyyy
Time commenced: |  |  |  |  | hh:mm
Displayed | Error shift limits |  | Sample info |  |  | Nominal values (%) |  | Sample stability | Correct for recov
name
GT | Min | Max | number | ID |  | PMB | moisture | Δ PMB | Yes/ No
1
2
3
Sample | Instrument | Error shift |  |  | RESULTS SUMMARY (Pass/fail)
number |  | Dry heat | Recovery | Corrected | Comments |  | Instrument | Type
1 | 1
2
3
1 | 2
2
3
Sample | Mean PMB (Start - ref) |  |  | Mean PMB (Dry H) |  | Mean PMB (Recovery)
number | Instr. 1 | Instr. 2 | Spare-ref | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | Spare-ref
1
2
3
Sample | PMB (Start - ref) |  |  | PMB (Dry heat) |  | PMB (Recovery) |  |  | Comments
number | Instr. 1 | Instr. 2 | Spare-ref | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | Spare-ref
1
2
3
Comments

## DampHeat
6.2.4 Damp heat [ref. OIML R xxx Annex C clause C.5.4]
Observer: | [5 empty] | General comments on test:
Type/ application #:
Instrument 1 ID:
7.2.4 Instrument 2 ID:
Spare instrument ID:
|  | Starf ref | Damp H | Recovery
EUT ambient temp: |  |  |  |  | °C
EUT ambient RH: |  |  |  |  | %
Sample temp: |  |  |  |  | °C
Spare ambient temp: |  |  |  |  | °C
Spare ambient RH: |  |  |  |  | %
Date commenced: |  |  |  |  | ddmmyyyy
Time commenced: |  |  |  |  | hh:mm
Displayed | Error shift limits |  | Sample info |  |  | Nominal values (%) |  | Sample stability | Correct for recov
name
GT | Min | Max | number | ID |  | PMB | moisture | Δ PMB | Yes/ No
1
2
3
Sample | Instrument | Error shift |  |  | RESULTS SUMMARY (Pass/fail)
number |  | Damp H | Recovery | Corrected | Comments |  | Instrument | Type
1 | 1
2
3
1 | 2
2
3
Sample | Mean PMB (Start - ref) |  |  | Mean PMB (Damp H) |  | Mean PMB (Recovery)
number | Instr. 1 | Instr. 2 | Spare-ref | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | Spare-ref
1
2
3
Sample | PMB (Start - ref) |  |  | PMB (Damp heat) |  | PMB (Recovery) |  |  | Comments
number | Instr. 1 | Instr. 2 | Spare-ref | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | Spare-ref
1
2
3
Comments

## VoltVariations
6.2.5 AC mains voltage variations [ref. OIML R xxx Annex C clause C.5.5]
Observer: | [6 empty] | Displayed | Error shift limits |  | Standard deviation | Sample info |  | Nominal values (%)
Type/ application #: | [6 empty] | name
Instrument 1 ID: | [6 empty] | GT | Min | Max | Max | ID |  | PMB | moisture
Instrument 2 ID: | [9 empty] | 0.1
|  | Start - ref | Influence |  | Recovery |  |  | General comments on test:
|  | Unom | Unom +10% | Unom -15% | Unom
Test voltage: | [5 empty] | V
Test frequency: | [5 empty] | Hz
Ambient temp: | [5 empty] | °C
Ambient RH: | [5 empty] | %
Date commenced: | [5 empty] | ddmmyy
Time commenced: | [5 empty] | hh:mm
Instrument | Standard deviation replicate PMB |  |  |  | Mean PMB |  |  |  | Error shift |  |  | RESULTS SUMMARY (Pass / fail)
| Start - ref | Unom +10% | Unom -15% | Recovery | Start - ref | Unom +10% | Unom -15% | Recovery | Unom +10% | Unom -15% | Recovery | Comments | Instrument | Type
1
2
Raw data entry
# | PMB (Start -ref) |  |  | PMB (Unom +10%) |  |  | PMB (Unom -15%) |  |  | PMB (Recovery)
| Instr. 1 | Instr. 2 | Comments | Instr. 1 | Instr. 2 | Comments | Instr. 1 | Instr. 2 | Comments | Instr. 1 | Instr. 2 | Comments
1
2
3
4
5
6
7
8
9
10

## BatteryV
6.2.6 Variation in voltage supplied by external 12V and 24 V road vehicle batteries [ref. OIML R xxx Annex C clause C.5.6]
Observer: | [6 empty] | Displayed | Error shift limits |  | Sample info |  | Nominal values (%)
Type/ application #: | [6 empty] | name
Instrument 1 ID: | [6 empty] | GT | Min | Max | ID |  | PMB | moisture
Instrument 2 ID:
|  | Start - ref | Influence |  | Recovery |  |  | General comments on test:
|  | Unom | Low V | High V | Unom
Test voltage: | [5 empty] | V
Test frequency: | [5 empty] | Hz
Ambient temp: | [5 empty] | °C
Ambient RH: | [5 empty] | %
Date commenced: | [5 empty] | ddmmyy
Time commenced: | [5 empty] | hh:mm
Instrument | Mean PMB |  |  |  | Error shift |  |  | RESULTS SUMMARY (Pass / fail)
| Start - ref | Low V | High V | Recovery | Low V | High V | Recovery | Comments | Instrument | Type
1
2
Raw data entry
# | PMB (Start -ref) |  |  | PMB (Low V) |  |  | PMB (High V) |  |  | PMB (Recovery)
| Instr. 1 | Instr. 2 | Comments | Instr. 1 | Instr. 2 | Comments | Instr. 1 | Instr. 2 | Comments | Instr. 1 | Instr. 2 | Comments
1
2
3
4
5
6

## VoltDipsInterrupt
7.3 Tests for disturbances
7.3.1 AC mains voltage dips and short interruptions [ref. OIML R xxx Annex C clause C.6.1]
Observer: | [8 empty] | Instrument 1 |  | Instrument 2
Type/ application #: | [8 empty] | Start testing | End testing | Start testing | End testing
Instrument 1 ID:
Instrument 2 ID: | [6 empty] | Test voltage: | [5 empty] | V
[7 empty] | Test frequency: | [5 empty] | Hz
Displayed | Fault limits |  | Sample info | Nominal values (%) |  |  | Ambient temp: | [5 empty] | °C
name | [6 empty] | Ambient RH: | [5 empty] | %
GT | Min | Max | ID | PMB | moisture |  | Date: | [5 empty] | ddmmyy
[7 empty] | Time: | [5 empty] | hh:mm
Mean ref PMB:
General comments on test: | [7 empty] | 1 |  | 1
[8 empty] | 2 |  | 2
[8 empty] | 3 |  | 3
[8 empty] | 4 |  | 4
[8 empty] | 5 |  | 5
[8 empty] | 6 |  | 6
A fault that exceeds the limit is insignificant if 'acted on' or exempted from the definition of a significant fault.
Instrument | All faults within limits (Yes / no) |  | All insignificant (Yes / no) |  | RESULTS SUMMARY (Pass / fail)
[5 empty] | Comments |  | Instrument | Type
1
2
Disturbance settings |  |  | Test: |  | A | B | C | D
|  |  | Nominal voltage | [5 empty] | V
|  |  | Voltage reduction: |  | 0 × Unom | 0 × Unom | 0.7 × Unom | 0 × Unom | Hz
|  |  | New voltage: |  | 0 | 0 | fill cell C27 | 0 | V
|  |  | Duration: |  | 0.5 | 1 | 25/30* | 250/300* | cycles
*These values are for 50 Hz and 60 Hz respectively
Raw data entry - AC mains voltage dips and short interruptions [ref. OIML R xxx Annex C clause C.6.1]
Instrument 1
Test |  | A |  | B |  | C |  | D
|  | PMB | Comments | PMB | Comments | PMB | Comments | PMB | Comments
1
2
3
4
5
6
7
8
9
10
*
1
2
3
4
5
6
7
8
9
10
*
All faults within limits
All faults insignificant
Comments
Continued - Raw data entry - AC mains voltage dips and short interruptions [ref. OIML R xxx Annex C clause C.6.1]
Instrument 2
Test |  | A |  | B |  | C |  | D
|  | PMB | Comments | PMB | Comments | PMB | Comments | PMB | Comments
1
2
3
4
5
6
7
8
9
10
*
1
2
3
4
5
6
7
8
9
10
*
All faults within limits
All faults insignificant
Comments
(e.g. action details)

## BurstsMains
6.3.2 Bursts (transients) on AC mains [ref. OIML R xxx Annex C clause C.6.2]
Observer: | [8 empty] | Instrument 1 |  | Instrument 2
Type/ application #: | [8 empty] | Start testing | End testing | Start testing | End testing
Instrument 1 ID:
Instrument 2 ID: | [7 empty] | Test voltage: |  |  |  |  | V
[8 empty] | Test frequency: |  |  |  |  | Hz
Displayed | Fault limits |  | Instrument | Sample info | Nominal values (%) |  |  | Ambient temp: |  |  |  |  | °C
name | [7 empty] | Ambient RH: |  |  |  |  | %
GT | Min | Max |  | ID | PMB | moisture |  | Date: |  |  |  |  | ddmmyy
|  |  | 1 |  |  |  |  | Time: |  |  |  |  | hh:mm
|  |  | 2* |  |  |  |  | Mean ref PMB:
*Use another sample/ batch for other instrument if 1st sample unfit due to overuse. | [7 empty] | 1 |  | 1
[8 empty] | 2 |  | 2
General comments on test: | [7 empty] | 3 |  | 3
[8 empty] | 4 |  | 4
[8 empty] | 5 |  | 5
[8 empty] | 6 |  | 6
A fault that exceeds the limit is insignificant if 'acted on' or exempted from the definition of a significant fault.
Instrument | All faults within limits (Yes / no) |  | All insignificant (Yes / no) |  | RESULTS SUMMARY (Pass / fail)
[5 empty] | Comments |  | Instrument | Type
1
2
Disturbance severity |  |  | Amplitude: |  | 1 | kV
|  |  | Repetition rate: |  | 5 | kHz
Raw data entry - Bursts (transients) on AC mains [ref. OIML R xxx Annex C clause C.6.2]
Instrument 1
Connection |  | L → G |  |  |  | N → G |  |  |  | PE → G
Polarity |  | ( + ) |  | ( - ) |  | ( + ) |  | ( - ) |  | ( + ) |  | ( - )
|  | PMB | Comments | PMB | Comments | PMB | Comments | PMB | Comments | PMB | Comments | PMB | Comments
1
2
3
4
5
6
7
8
9
10
*
1
2
3
4
5
6
7
8
9
10
*
All faults within limits
All faults insignificant
Comments
Continued - Raw data entry - Bursts (transients) on AC mains [ref. OIML R xxx Annex C clause C.6.2]
Instrument 2
Connection |  | L → G |  |  |  | N → G |  |  |  | PE → G
Polarity |  | ( + ) |  | ( - ) |  | ( + ) |  | ( - ) |  | ( + ) |  | ( - )
|  | PMB | Comments | PMB | Comments | PMB | Comments | PMB | Comments | PMB | Comments | PMB | Comments
1
2
3
4
5
6
7
8
9
10
*
1
2
3
4
5
6
7
8
9
10
*
All faults within limits
All faults insignificant
Comments
(e.g. action details)

## RadiatedRF,EMFields
6.3.3 Radiated, radio-frequency, electromagnetic fields [ref. OIML R xxx Annex C clause C.6.3]
Observer: | [5 empty] | Displayed | Fault limits
Type/ application #: | [5 empty] | name
Instrument 1 ID: | [5 empty] | GT | Min | Max
Instrument 2 ID:
General comments on test:
Disturbance settings (indicate values if alternative settings are applied)
Frequency range: |  |  | MHz |  | Frequency range: |  | 26* - 2000 | MHz
Modulation: |  |  |  |  | Modulation: |  | 80% AM, 1 kHz sine wave
Field strength: |  |  | V/ m |  | Field strength: |  | 10 | V/ m
Step size: |  |  |  |  | Step size: |  | 1 | %
Rate of sweep: |  |  |  |  | Rate of sweep: |  | 0.0015 | decade/ s
*Testing from 80 MHz is permitted.
Antenna:
Circular polarisation (yes / no):
Note: Two orthogonal polarisations shall be tested if the answer is 'No'.
Example radiated EMS testing scheme:
| Antenna polarisation: | 1 - Vertical |  |  |  | 2 - Horizontal
| Facing EUT side: | Front | Left | Right | Back | Front | Left | Right | Back
A fault that exceeds the limit is insignificant if 'acted on' or exempted from the definition of a significant fault.
Instrument | All faults within limits (Yes / no) | 1 - Vertical Left: All insignificant (Yes / no); 1 - Vertical Back: RESULTS SUMMARY (Pass / fail)
[5 empty] | Comments |  | Instrument | Type
1
2
Raw data entry - Radiated, radio-frequency, electromagnetic fields
[ref. OIML R xxx Annex C clause C.6.3]
Instrument number:
Before disturbances |  | 1 - Vertical Right: During disturbances
| Antenna polarisation: |  |  | Frequency (MHz) | PMB | Fault | Within limit (Yes / no) | Insignificant (Yes / no) | Comments
Facing EUT side:
Test voltage:
Test frequency:
Ambient temp:
Ambient RH:
Date:
Time:
Grain sample ID:
Nominal PMB:
Nominal moisture:
Mean ref PMB:
1
2
3
4
5
6
Use another sample/ batch for other
orientations or instrument if 1st
sample unfit due to overuse/ exposure.
Comments:
*
Add additional rows if required.
Continued - Raw data entry - Radiated, radio-frequency, electromagnetic fields
[ref. OIML R xxx Annex C clause C.6.3]
Instrument number:
Before disturbances |  | 1 - Vertical Right: During disturbances
| Antenna polarisation: |  |  | Frequency (MHz) | PMB | Fault | Within limit (Yes / no) | Insignificant (Yes / no) | Comments
Facing EUT side:
Test voltage:
Test frequency:
Ambient temp:
Ambient RH:
Date:
Time:
Grain sample ID:
Nominal PMB:
Nominal moisture:
Mean ref PMB:
1
2
3
4
5
6
Use another sample/ batch for other
orientations or instrument if 1st
sample unfit due to overuse/ exposure.
Comments:
*
Add additional rows if required.
Continued - Raw data entry - Radiated, radio-frequency, electromagnetic fields
[ref. OIML R xxx Annex C clause C.6.3]
Instrument number:
Before disturbances |  | 1 - Vertical Right: During disturbances
| Antenna polarisation: |  |  | Frequency (MHz) | PMB | Fault | Within limit (Yes / no) | Insignificant (Yes / no) | Comments
Facing EUT side:
Test voltage:
Test frequency:
Ambient temp:
Ambient RH:
Date:
Time:
Grain sample ID:
Nominal PMB:
Nominal moisture:
Mean ref PMB:
1
2
3
4
5
6
Use another sample/ batch for other
orientations or instrument if 1st
sample unfit due to overuse/ exposure.
Comments:
*
Add additional rows if required.
Continued - Raw data entry - Radiated, radio-frequency, electromagnetic fields
[ref. OIML R xxx Annex C clause C.6.3]
Instrument number:
Before disturbances |  | 1 - Vertical Right: During disturbances
| Antenna polarisation: |  |  | Frequency (MHz) | PMB | Fault | Within limit (Yes / no) | Insignificant (Yes / no) | Comments
Facing EUT side:
Test voltage:
Test frequency:
Ambient temp:
Ambient RH:
Date:
Time:
Grain sample ID:
Nominal PMB:
Nominal moisture:
Mean ref PMB:
1
2
3
4
5
6
Use another sample/ batch for other
orientations or instrument if 1st
sample unfit due to overuse/ exposure.
Comments:
*
Add additional rows if required.
Raw data entry - Radiated, radio-frequency, electromagnetic fields
[ref. OIML R xxx Annex C clause C.6.3]
Instrument number:
Before disturbances |  | 1 - Vertical Right: During disturbances
| Antenna polarisation: |  |  | Frequency (MHz) | PMB | Fault | Within limit (Yes / no) | Insignificant (Yes / no) | Comments
Facing EUT side:
Test voltage:
Test frequency:
Ambient temp:
Ambient RH:
Date:
Time:
Grain sample ID:
Nominal PMB:
Nominal moisture:
Mean ref PMB:
1
2
3
4
5
6
Use another sample/ batch for other
orientations or instrument if 1st
sample unfit due to overuse/ exposure.
Comments:
*
Add additional rows if required.
Continued - Raw data entry - Radiated, radio-frequency, electromagnetic fields
[ref. OIML R xxx Annex C clause C.6.3]
Instrument number:
Before disturbances |  | 1 - Vertical Right: During disturbances
| Antenna polarisation: |  |  | Frequency (MHz) | PMB | Fault | Within limit (Yes / no) | Insignificant (Yes / no) | Comments
Facing EUT side:
Test voltage:
Test frequency:
Ambient temp:
Ambient RH:
Date:
Time:
Grain sample ID:
Nominal PMB:
Nominal moisture:
Mean ref PMB:
1
2
3
4
5
6
Use another sample/ batch for other
orientations or instrument if 1st
sample unfit due to overuse/ exposure.
Comments:
*
Add additional rows if required.
Continued - Raw data entry - Radiated, radio-frequency, electromagnetic fields
[ref. OIML R xxx Annex C clause C.6.3]
Instrument number:
Before disturbances |  | 1 - Vertical Right: During disturbances
| Antenna polarisation: |  |  | Frequency (MHz) | PMB | Fault | Within limit (Yes / no) | Insignificant (Yes / no) | Comments
Facing EUT side:
Test voltage:
Test frequency:
Ambient temp:
Ambient RH:
Date:
Time:
Grain sample ID:
Nominal PMB:
Nominal moisture:
Mean ref PMB:
1
2
3
4
5
6
Use another sample/ batch for other
orientations or instrument if 1st
sample unfit due to overuse/ exposure.
Comments:
*
Add additional rows if required.
Continued - Raw data entry - Radiated, radio-frequency, electromagnetic fields
[ref. OIML R xxx Annex C clause C.6.3]
Instrument number:
Before disturbances |  | 1 - Vertical Right: During disturbances
| Antenna polarisation: |  |  | Frequency (MHz) | PMB | Fault | Within limit (Yes / no) | Insignificant (Yes / no) | Comments
Facing EUT side:
Test voltage:
Test frequency:
Ambient temp:
Ambient RH:
Date:
Time:
Grain sample ID:
Nominal PMB:
Nominal moisture:
Mean ref PMB:
1
2
3
4
5
6
Use another sample/ batch for other
orientations or instrument if 1st
sample unfit due to overuse/ exposure.
Comments:
*
Add additional rows if required.

## ConductedRF,EMFields
6.3.4 Conducted, radio-frequency, electromagnetic fields [ref. OIML R xxx Annex C clause C.6.5]
Observer: | [9 empty] | Instrument 1 |  | Instrument 2
Type/ application #: | [9 empty] | Start testing | End testing | Start testing | End testing
Instrument 1 ID:
Instrument 2 ID: | [8 empty] | Test voltage: |  |  |  |  | V
[9 empty] | Test frequency: |  |  |  |  | Hz
Displayed | Fault limits |  | Instrument | Sample info | Nominal values (%) |  |  |  | Ambient temp: |  |  |  |  | °C
name | [8 empty] | Ambient RH: |  |  |  |  | %
GT | Min | Max |  | ID | PMB | moisture |  |  | Date: |  |  |  |  | ddmmyy
|  |  | 1 | [5 empty] | Time: |  |  |  |  | hh:mm
|  |  | 2* | [5 empty] | Mean ref PMB:
*Use another sample/ batch for other instrument if 1st sample unfit due to overuse. | [8 empty] | 1 |  | 1
[9 empty] | 2 |  | 2
General comments on test: | [8 empty] | 3 |  | 3
[9 empty] | 4 |  | 4
[9 empty] | 5 |  | 5
[9 empty] | 6 |  | 6
A fault that exceeds the limit is insignificant if 'acted on' or exempted from the definition of a significant fault.
Instrument | All faults within limit (Yes / no) |  | All insignificant (Yes / no) |  | RESULTS SUMMARY (Pass / fail)
[5 empty] | Comments |  | Instrument | Type
1
2
Disturbance settings (Indicate values if alternative settings are applied)
Frequency range: |  |  | MHz |  |  |  |  | Frequency range: |  | 0.15 - 80 | MHz
Modulation: | [7 empty] | Modulation: |  | 80% AM, 1 kHz sine wave
RF amplitude (50 Ω): |  |  | V (e.m.f.) |  |  |  |  | RF amplitude (50 Ω): |  | 10 | V (e.m.f.)
Raw data entry - Radiated, radio-frequency, electromagnetic fields [ref. OIML R xxx Annex C clause C.6.3]
Instrument 1 | [7 empty] | Instrument 2
Cable or interface: | [7 empty] | Cable or interface:
Frequency (MHz) | PMB | Fault | Within limit (Yes / no) | Insignificant (Yes / no) | Comments |  |  | Frequency (MHz) | PMB | Fault | Within limit (Yes / no) | Insignificant (Yes / no) | Comments
* | [7 empty] | *
Add additional rows if required.
Continued - Instrument 1 | [7 empty] | Continued - Instrument 2
Cable or interface: | [7 empty] | Cable or interface:
Frequency (MHz) | PMB | Fault | Within limit (Yes / no) | Insignificant (Yes / no) | Comments |  |  | Frequency (MHz) | PMB | Fault | Within limit (Yes / no) | Insignificant (Yes / no) | Comments
* | [7 empty] | *
Add additional rows if required.
Continued - Raw data entry - Radiated, radio-frequency, electromagnetic fields [ref. OIML R xxx Annex C clause C.6.3]
Continued - Instrument 1 | [7 empty] | Continued - Instrument 2
Cable or interface: | [7 empty] | Cable or interface:
Frequency (MHz) | PMB | Fault | Within limit (Yes / no) | Insignificant (Yes / no) | Comments |  |  | Frequency (MHz) | PMB | Fault | Within limit (Yes / no) | Insignificant (Yes / no) | Comments
* | [7 empty] | *
Add additional rows if required.
Continued - Instrument 1 | [7 empty] | Continued - Instrument 2
Cable or interface: | [7 empty] | Cable or interface:
Frequency (MHz) | PMB | Fault | Within limit (Yes / no) | Insignificant (Yes / no) | Comments |  |  | Frequency (MHz) | PMB | Fault | Within limit (Yes / no) | Insignificant (Yes / no) | Comments
* | [7 empty] | *
Add additional rows if required.
Copy table to additional pages ir required.

## ElecDischarges
6.3.5 Electrostatic discharges [ref. OIML R xxx Annex C clause C.6.5]
Observer: | [8 empty] | Instrument 1 |  | Instrument 2
Type/ application #: | [8 empty] | Start testing | End testing | Start testing | End testing
Instrument 1 ID:
Instrument 2 ID: | [7 empty] | Test voltage: |  |  |  |  | V
[8 empty] | Test frequency: |  |  |  |  | Hz
Displayed | Fault limits |  | Instrument | Sample info | Nominal values (%) |  |  | Ambient temp: |  |  |  |  | °C
name | [7 empty] | Ambient RH: |  |  |  |  | %
GT | Min | Max |  | ID | PMB | moisture |  | Date: |  |  |  |  | ddmmyy
|  |  | 1 |  |  |  |  | Time: |  |  |  |  | hh:mm
|  |  | 2* |  |  |  |  | Mean ref PMB:
*Use another sample/ batch for other instrument if 1st sample unfit due to overuse. | [7 empty] | 1 |  | 1
[8 empty] | 2 |  | 2
General comments on test: | [7 empty] | 3 |  | 3
[8 empty] | 4 |  | 4
[8 empty] | 5 |  | 5
[8 empty] | 6 |  | 6
A fault that exceeds the limit is insignificant if 'acted on' or exempted from the definition of a significant fault.
Instrument | All faults within limits (Yes / no) |  | All insignificant (Yes / no) |  | RESULTS SUMMARY (Pass / fail)
[5 empty] | Comments |  | Instrument | Type
1
2
Direct discharge mode (contact, paint penetration or air) | [6 empty] | Notes: H = horizontal, V = vertical
2 kV | [6 empty] | Refer to diagram of coupling plane positions for applying discharges indirectly.
4 kV
6 kV | [8 empty] | V2
8 kV
[7 empty] | V1 |  | EUT |  | V3
It is permitted to only test one at polarity. | [8 empty] | front
IEC 61000-4-2 specifies to select the most sensitive polarity.
V4
Polarity of indirect discharges (positive / negative):
Raw data entry - Electrostatic discharges [ref. OIML R xxx Annex C clause C.6.5]
Instrument 1 - PMB - Direct discharge application
Voltage (kV) |  | 2 |  |  | 4 |  |  | 6 |  |  | 8
Polarity |  | ( + ) | ( - ) | Comments | ( + ) | ( - ) | Comments | ( + ) | ( - ) | Comments | ( + ) | ( - ) | Comments
1
2
3
4
5
6
7
8
9
10
*
1
2
3
4
5
6
7
8
9
10
*
All faults within limits
All faults insignificant
Comments
Continued - Raw data entry - Electrostatic discharges [ref. OIML R xxx Annex C clause C.6.5]
Instrument 1 - PMB - Indirect discharge application
Voltage (kV) |  | V1 |  |  |  | V2 |  |  |  | V3
Polarity |  | 2 | 4 | 6 | Comments | 2 | 4 | 6 | Comments | 2 | 4 | 6 | Comments
1
2
3
4
5
6
7
8
9
10
*
1
2
3
4
5
6
7
8
9
10
*
All faults within limits
All faults insignificant
Comments
(e.g. action details)
Continued - Raw data entry - Electrostatic discharges [ref. OIML R xxx Annex C clause C.6.5]
Continued - Instrument 1 - PMB - Indirect discharge application
Voltage (kV) |  | V3 |  |  |  | H
Polarity |  | 2 | 4 | 6 | Comments | 2 | 4 | 6 | Comments
1
2
3
4
5
6
7
8
9
10
*
1
2
3
4
5
6
7
8
9
10
*
All faults within limits
All faults insignificant
Comments
(e.g. action details)
Continued - Raw data entry - Electrostatic discharges [ref. OIML R xxx Annex C clause C.6.5]
Instrument 2 - PMB - Direct discharge application
Voltage (kV) |  | 2 |  |  | 4 |  |  | 6 |  |  | 8
Polarity |  | ( + ) | ( - ) | Comments | ( + ) | ( - ) | Comments | ( + ) | ( - ) | Comments | ( + ) | ( - ) | Comments
1
2
3
4
5
6
7
8
9
10
*
1
2
3
4
5
6
7
8
9
10
*
All faults within limits
All faults insignificant
Comments
(e.g. action details)
Continued - Raw data entry - Electrostatic discharges [ref. OIML R xxx Annex C clause C.6.5]
Instrument 2 - PMB - Indirect discharge application
Voltage (kV) |  | V1 |  |  |  | V2 |  |  |  | V3
Polarity |  | 2 | 4 | 6 | Comments | 2 | 4 | 6 | Comments | 2 | 4 | 6 | Comments
1
2
3
4
5
6
7
8
9
10
*
1
2
3
4
5
6
7
8
9
10
*
All faults within limits
All faults insignificant
Comments
(e.g. action details)
Continued - Raw data entry - Electrostatic discharges [ref. OIML R xxx Annex C clause C.6.5]
Continued - Instrument 2 - PMB - Indirect discharge application
Voltage (kV) |  | V3 |  |  |  | H
Polarity |  | 2 | 4 | 6 | Comments | 2 | 4 | 6 | Comments
1
2
3
4
5
6
7
8
9
10
*
1
2
3
4
5
6
7
8
9
10
*
All faults within limits
All faults insignificant
Comments
(e.g. action details)

## StorageTemp
6.3.6 Storage temperature (extreme shipping conditions) [ref. OIML R xxx Annex C clause C.6.6]
Observer: | [5 empty] | General comments on test:
Type/ application #:
Instrument 1 ID:
Instrument 2 ID:
|  | Start testing | End testing
Ambient temp: |  |  |  | °C
Ambient RH: |  |  |  | %
Date: |  |  |  | ddmmyy
Time: |  |  |  | hh:mm
Displayed | Fault limits |  | Sample info | Nominal values (%)
name
GT | Min | Max | ID | PMB | moisture
Instrument | All faults within limit (Yes / no) |  | All insignificant (Yes / no) |  | RESULTS SUMMARY (Pass / fail)
[5 empty] | Comments |  | Instrument | Type
1
2
Disturbance severity |  |  | Minimum temp: |  | -20 | °C
|  |  | Maximum temp: |  | 50 | °C
Raw data entry
BEFORE disturbance
# | PMB
| Instr. 1 | Instr. 2
1
2
3
4
5
6
7
8
9
10
Mean PMB
A fault that exceeds the limit is insignificant if 'acted on' or exempted from the definition of a significant fault.
AFTER disturbance
| Instrument 1 |  |  |  |  | Instrument 2
| PMB | Fault | Within limit (Yes / no) | Insignificant (Yes / no) |  | PMB | Fault | Within limit (Yes / no) | Insignificant (Yes / no)
1 |  |  |  |  | 1
2 |  |  |  |  | 2
3 |  |  |  |  | 3
4 |  |  |  |  | 4
5 |  |  |  |  | 5
6 |  |  |  |  | 6
7 |  |  |  |  | 7
8 |  |  |  |  | 8
9 |  |  |  |  | 9
10 |  |  |  |  | 10
| Comments / details of action |  |  |  |  | Comments / details of action

## RandVibration
6.3.7 Random vibrations [ref. OIML R xxx Annex C clause C.6.7]
Observer: | [5 empty] | General comments on test:
Type/ application #:
Instrument 1 ID:
Instrument 2 ID:
|  | Start testing | End testing
Ambient temp: |  |  |  | °C
Ambient RH: |  |  |  | %
Date: |  |  |  | ddmmyy
Time: |  |  |  | hh:mm
Displayed | Fault limits |  | Sample info | Nominal values (%)
name
GT | Min | Max | ID | PMB | moisture
Instrument | All faults within limits (Yes / no) |  | All insignificant (Yes / no) |  | RESULTS SUMMARY (Pass / fail)
[5 empty] | Comments |  | Instrument | Type
1
2
Disturbance severity |  |  | Total frequency range: |  |  | 10 - 150 | Hz
|  |  | Total RMS level: |  |  | 7 | ms-2
|  |  | ASD level 10 Hz – 20 Hz: |  |  | 1 | m2s-3
|  |  | ASD level 20 Hz – 150 Hz |  |  | -3 | dB/octave
|  |  | Number of directions: |  |  | 3
|  |  | Duration per axis: |  |  | 60 | minutes
Raw data entry
BEFORE disturbance
# | PMB
| Instr. 1 | Instr. 2
1
2
3
4
5
6
Mean PMB
A fault that exceeds the limit is insignificant if 'acted on' or exempted from the definition of a significant fault.
AFTER disturbance
| Instrument 1 |  |  |  |  | Instrument 2
| PMB | Fault | Within limit (Yes / no) | Insignificant (Yes / no) |  | PMB | Fault | Within limit (Yes / no) | Insignificant (Yes / no)
1 |  |  |  |  | 1
2 |  |  |  |  | 2
3 |  |  |  |  | 3
4 |  |  |  |  | 4
5 |  |  |  |  | 5
6 |  |  |  |  | 6
7 |  |  |  |  | 7
8 |  |  |  |  | 8
9 |  |  |  |  | 9
10 |  |  |  |  | 10
| Comments / details of action |  |  |  |  | Comments / details of action

## Accuracy,r&R-Calibration
6.4 Tests to assess the submitted calibrations
6.4.1 Accuracy and precision [ref. OIML R xxx Annex C clause C.7.1]
Observer: | [5 empty] | General comments on test settings:
Type/ application #:
Instrument 1 ID:
Instrument 2 ID:
|  | GT1 |  | GT2 |  | GT3 |  | GT4 |  | GT*
|  | Start ref | End ref | Start ref | End ref | Start ref | End ref | Start ref | End ref | Start ref | End ref
Ambient temp: | [11 empty] | °C
Ambient RH: | [11 empty] | %
Date commenced: | [11 empty] | ddmmyyyy
Time commenced: | [11 empty] | hh:mm
Submitted calibrations | [5 empty] | RESULTS SUMMARY (Pass/fail)
Number | Calibration name |  | Version | Disp. name | MB (%) | Accuracy | Repeat. | Reprod. | Comments
GT1
GT2
GT3
GT4
GT*
Limits
Calibration | y(bar) |  | SEP | Pooled SD | SDDI
number | Min | Max | Max | Max | Max
GT1
GT2
GT3
GT4
GT*
* Add cells for additional calibrations if necessary
Calibration | Sample info |  | Reference values |  | Sample info |  | Reference values |  | Sample info |  | Reference values
number | number | ID | PMB | moisture | number | ID | PMB | moisture | number | ID | PMB | moisture
GT1 | 1 |  |  |  | 11 |  |  |  | 21
| 2 |  |  |  | 12 |  |  |  | 22
| 3 |  |  |  | 13 |  |  |  | 23
| 4 |  |  |  | 14 |  |  |  | 24
| 5 |  |  |  | 15 |  |  |  | 25
| 6 |  |  |  | 16 |  |  |  | 26
| 7 |  |  |  | 17 |  |  |  | 27
| 8 |  |  |  | 18 |  |  |  | 28
| 9 |  |  |  | 19 |  |  |  | 29
| 10 |  |  |  | 20 |  |  |  | 30*
GT2 | 1 |  |  |  | 11 |  |  |  | 21
| 2 |  |  |  | 12 |  |  |  | 22
| 3 |  |  |  | 13 |  |  |  | 23
| 4 |  |  |  | 14 |  |  |  | 24
| 5 |  |  |  | 15 |  |  |  | 25
| 6 |  |  |  | 16 |  |  |  | 26
| 7 |  |  |  | 17 |  |  |  | 27
| 8 |  |  |  | 18 |  |  |  | 28
| 9 |  |  |  | 19 |  |  |  | 29
| 10 |  |  |  | 20 |  |  |  | 30*
* Add cells for additional samples if necessary
Calibration | Sample info |  | Reference values |  | Sample info |  | Reference values |  | Sample info |  | Reference values
number | number | ID | PMB | moisture | number | ID | PMB | moisture | number | ID | PMB | moisture
GT3 | 1 |  |  |  | 11 |  |  |  | 21
| 2 |  |  |  | 12 |  |  |  | 22
| 3 |  |  |  | 13 |  |  |  | 23
| 4 |  |  |  | 14 |  |  |  | 24
| 5 |  |  |  | 15 |  |  |  | 25
| 6 |  |  |  | 16 |  |  |  | 26
| 7 |  |  |  | 17 |  |  |  | 27
| 8 |  |  |  | 18 |  |  |  | 28
| 9 |  |  |  | 19 |  |  |  | 29
| 10 |  |  |  | 20 |  |  |  | 30*
GT4 | 1 |  |  |  | 11 |  |  |  | 21
| 2 |  |  |  | 12 |  |  |  | 22
| 3 |  |  |  | 13 |  |  |  | 23
| 4 |  |  |  | 14 |  |  |  | 24
| 5 |  |  |  | 15 |  |  |  | 25
| 6 |  |  |  | 16 |  |  |  | 26
| 7 |  |  |  | 17 |  |  |  | 27
| 8 |  |  |  | 18 |  |  |  | 28
| 9 |  |  |  | 19 |  |  |  | 29
| 10 |  |  |  | 20 |  |  |  | 30*
GT*
* Add cells for additional calibrations and/or samples if necessary
[9 empty] | GT1 | on Instr.1 | GT1 | on Instr. 2
The requirement for y(bar) was fulfilled and all functions operated as designed (Pass or Fail):
Calibration | Sample | Mean PMB |  |  | Error (yi) |  | Pooled yi, i.e. y(bar)
number | number | reference | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | MinLimit | MaxLimit
GT1 | 1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30*
Add cells for additional samples if necessary
[9 empty] | GT2 | on Instr.1 | GT2 | on Instr. 2
The requirement for y(bar) was fulfilled and all functions operated as designed (Pass or Fail):
Calibration | Sample | Mean PMB |  |  | Error (yi) |  | Pooled yi, i.e. y(bar)
number | number | reference | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | MinLimit | MaxLimit
GT2 | 1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30*
GT*
* Add cells for additional calibrations and/or samples if necessary
[9 empty] | GT1 | on Instr.1 | GT1 | on Instr. 2
The requirement for SEP was fulfilled and all functions operated as designed (Pass or Fail):
Calibration | Sample | Reference | PMB (j=1) |  | Error (y,j=1) |  | SEP
number | number | PMB | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | MaxLimit
GT1 | 1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30*
* Add cells for additional samples if necessary
[9 empty] | GT2 | on Instr.1 | GT2 | on Instr. 2
The requirement for SEP was fulfilled and all functions operated as designed (Pass or Fail):
Calibration | Sample | Reference | PMB (j=1) |  | Error (y,j=1) |  | SEP
number | number | PMB | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | MaxLimit
GT2 | 1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30*
GT*
* Add cells for additional calibrations and/or samples if necessary
[9 empty] | GT1 | on Instr.1 | GT1 | on Instr. 2
The requirement for repeatability was fulfilled and all functions operated as designed (Pass or Fail):
Calibration | Sample | SD |  | SD2 (variance) |  | Pooled SD
number | number | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | MaxLimit
GT1 | 1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30*
* Add cells for additional samples if necessary
[9 empty] | GT2 | on Instr.1 | GT2 | on Instr. 2
The requirement for repeatability was fulfilled and all functions operated as designed (Pass or Fail):
Calibration | Sample | SD |  | SD2 (variance) |  | Pooled SD
number | number | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | Instr. 1 | Instr. 2 | MaxLimit
GT2 | 1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30*
GT*
* Add cells for additional calibrations and/or samples if necessary
[9 empty] | GT1 | on Instr.1 | GT1 | on Instr. 2
The requirement for reproducibility was fulfilled and all functions operated as designed (Pass or Fail):
Calibration | Sample | Mean PMB |  | d | SDDI | SSDI
number | number | Instr. 1 | Instr. 2 |  |  | MaxLimit
GT1 | 1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30*
* Add cells for additional samples if necessary
[9 empty] | GT2 | on Instr.1 | GT2 | on Instr. 2
The requirement for reproducibility was fulfilled and all functions operated as designed (Pass or Fail):
Calibration | Sample | Mean PMB |  | d | SDDI | SSDI
number | number | Instr. 1 | Instr. 2 |  |  | MaxLimit
GT2 | 1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
GT*
* Add cells for additional calibrations and/or samples if necessary
Raw data entry - Accuracy and precision [ref. OIML R xxx Annex C clause C.7.1]
Calibration | Sample | PMB @ ref |  | Comments | Sample | PMB @ ref |  | Comments | Sample | PMB @ ref |  | Comments
number | number | Instr. 1 | Instr.2 |  | number | Instr. 1 | Instr.2 |  | number | Instr. 1 | Instr.2
GT1 | 1 |  |  |  | 11 |  |  |  | 21
| 2 |  |  |  | 12 |  |  |  | 22
| 3 |  |  |  | 13 |  |  |  | 23
| 4 |  |  |  | 14 |  |  |  | 24
| 5 |  |  |  | 15 |  |  |  | 25
| 6 |  |  |  | 16 |  |  |  | 26
| 7 |  |  |  | 17 |  |  |  | 27
| 8 |  |  |  | 18 |  |  |  | 28
| 9 |  |  |  | 19 |  |  |  | 29
| 10 |  |  |  | 20 |  |  |  | 30*
* Add cells for additional samples if necessary
Continued - Raw data entry - Accuracy and precision [ref. OIML R xxx Annex C clause C.7.1]
Calibration | Sample | PMB @ ref |  | Comments | Sample | PMB @ ref |  | Comments | Sample | PMB @ ref |  | Comments
number | number | Instr. 1 | Instr.2 |  | number | Instr. 1 | Instr.2 |  | number | Instr. 1 | Instr.2
GT2 | 1 |  |  |  | 11 |  |  |  | 21
| 2 |  |  |  | 12 |  |  |  | 22
| 3 |  |  |  | 13 |  |  |  | 23
| 4 |  |  |  | 14 |  |  |  | 24
| 5 |  |  |  | 15 |  |  |  | 25
| 6 |  |  |  | 16 |  |  |  | 26
| 7 |  |  |  | 17 |  |  |  | 27
| 8 |  |  |  | 18 |  |  |  | 28
| 9 |  |  |  | 19 |  |  |  | 29
| 10 |  |  |  | 20 |  |  |  | 30*
GT*
* Add cells for additional calibrations and/or samples if necessary
Note on G260: Leida Queddeng (NMIA): SQRT(sum variances / n) (min n = 30)
Note on H260: Leida Queddeng (NMIA): SQRT(sum variances / n) (min n = 30)
Note on G296: Leida Queddeng (NMIA): SQRT(sum variances / n) (min n = 30)
Note on H296: Leida Queddeng (NMIA): SQRT(sum variances / n) (min n = 30)

## STS-Calibration
6.4.2 Sample temperature sensitivity (STS) [OIML Rxxx-2, clause C.7.2]
Observer: | [8 empty] | Start ref | Tref - ΔTC | Recov 1 | Tref + ΔTH | Recov 2
Type/ application #: | [6 empty] | Ambient temp: | [6 empty] | °C
Instrument 1 ID: | [6 empty] | Ambient RH: | [6 empty] | %
Instrument 2 ID: | [6 empty] | Sample temp: | [6 empty] | °C
[7 empty] | Date commenced: | [6 empty] | ddmmyyyy
[7 empty] | Time commenced: | [6 empty] | hh:mm
General comments on test settings:
Limits
Calibrations submitted and tested | [6 empty] | Pooled error shift |  | RESULTS SUMMARY
number | Calibration name |  | Version | Displayed name |  | MB (%) | Min | Max | Pass/fail | Comments
GT1
GT2
GT3
GT4
Calibration | Sample info |  |  | Nomimal values (%) |  |  |  | Calibration | Sample info |  |  | Nomimal values (%)
number | moisture | number | ID | PMB | moisture |  |  | number | moisture | number | ID | PMB | moisture
GT1 | low m | 1 | [5 empty] | GT3 | low m | 1
|  | 2 | [7 empty] | 2
|  | 3 | [7 empty] | 3
| high m | 4 | [6 empty] | high m | 4
|  | 5 | [7 empty] | 5
|  | 6 | [7 empty] | 6
GT2 | high m | 1 | [5 empty] | GT4 | high m | 1
|  | 2 | [7 empty] | 2
|  | 3 | [7 empty] | 3
| low m | 4 | [6 empty] | low m | 4
|  | 5 | [7 empty] | 5
|  | 6 | [7 empty] | 6
Instrument 1 only
Calibration | Sample info |  | Mean PMB |  |  |  |  | Error shift |  | Pooled error shift |  | RESULT | Limits pooled e/shift
number | moisture | number | Start ref | Tref - ΔTC | Recov 1 | Tref + ΔTH | Recov 2 | Tref - ΔTC | Tref + ΔTH | Tref - ΔTC | Tref + ΔTH | Pass/fail | Min | Max
GT1 | low m | 1
2
3
| high m | 4
5
6
GT2 | high m | 1
2
3
| low m | 4
5
6
GT3 | low m | 1
2
3
| high m | 4
5
6
GT4 | high m | 1
2
3
| low m | 4
5
6
[8 empty] | All functions operated as intended: |  |  |  |  | Pass/fail
Instrument 2 only
Calibration | Sample info |  | Mean PMB |  |  |  |  | Error shift |  | Pooled error shift |  | RESULT | Limits pooled e/shift
number | moisture | number | Start ref | Tref - ΔTC | Recov 1 | Tref + ΔTH | Recov 2 | Tref - ΔTC | Tref + ΔTH | Tref - ΔTC | Tref + ΔTH | Pass/fail | Min | Max
GT1 | low m | 1
2
3
| high m | 4
5
6
GT2 | high m | 1
2
3
| low m | 4
5
6
GT3 | low m | 1
2
3
| high m | 4
5
6
GT4 | high m | 1
2
3
| low m | 4
5
6
[8 empty] | All functions operated as intended: |  |  |  |  | Pass/fail
Raw data entry - Sample temperature sensitivity (STS) [ref. OIML R xxx Annex C clause C.7.2]
Calibration | Sample info |  | PMB (start ref) |  | PMB (Tref - ΔTC) |  | PMB (recov 1) |  | PMB (Tref + ΔTH) |  | PMB (recov 2) |  | Comments e.g. functionality
number | moisture | number | Instr. 1 | Instr.2 | Instr. 1 | Instr.2 | Instr. 1 | Instr.2 | Instr. 1 | Instr.2 | Instr. 1 | Instr.2
GT1 | low m | 1
2
3
| high m | 4
5
6
GT2 | high m | 1
2
3
| low m | 4
5
6
*
* Add cells for additional calibrations if necessary
Continued - Raw data entry - Sample temperature sensitivity (STS) [ref. OIML R xxx Annex C clause C.7.2]
Calibration | Sample info |  | PMB (start ref) |  | PMB (Tref - ΔTC) |  | PMB (recov 1) |  | PMB (Tref + ΔTH) |  | PMB (recov 2) |  | Comments e.g. functionality
number | moisture | number | Instr. 1 | Instr.2 | Instr. 1 | Instr.2 | Instr. 1 | Instr.2 | Instr. 1 | Instr.2 | Instr. 1 | Instr.2
GT3 | low m | 1
2
3
| high m | 4
5
6
GT4 | high m | 1
2
3
| low m | 4
5
6

## FormSelections
NAME | PassOrFail | PassFailNA | YesOrNo | YesNoNA
| Pass | Pass | Yes | Yes
| Fail | Fail | No | No
|  | NA |  | NA